## Slide 1
365 nm UV-A AUTOFLUORESCENCE
850 nm NIR ABSORPTION
ARDUINO UNO R4 EMBEDDED
OPENCV COMPUTER VISION
BloodTrace™
Dual-Wavelength Optical Contamination Screening Platform
Rapid, Non-Destructive Bioburden Localization for Reusable Surgical Instruments Prior to Autoclave Sterilization
PRESENTER & PROJECT LEAD
Devansh Lila
Department of Mechanical / Biomedical Engineering
Malaviya National Institute of Technology (MNIT) Jaipur
KEY INNOVATION METRICS
• Hardware Bill-of-Materials: ~$64.00 (Low-resource healthcare accessibility)
• Screening Duration: < 3 seconds per instrument (Pre-autoclave quality gate)
• Multi-Spectral Physics: 365nm fluorophore emission + 850nm heme absorption
⚠ REGULATORY SAFETY NOTICE & PROTOTYPE DEMARCATION
BloodTrace is an experimental college-level engineering proof-of-concept screening prototype. It is NOT an FDA-cleared in-vitro diagnostic device. Categorizes candidate regions as 'Suspected Residue' to assist human inspection; cannot establish sterility or release tools for surgery.

## Slide 2
CLINICAL PROBLEM & MOTIVATION
Why Residual Soil on Reusable Instruments Demands Objective Screening
02 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
Bioburden In Crevices
Blood, cellular debris, and proteins become mechanically trapped in serrated jaws, box locks, and hinged joints of surgical forceps and clamps during invasive surgery. Incomplete manual washing leaves microscopic residues in these recesses.
The 'Baking' Hazard
Autoclave sterilization relies on direct pressurized saturated steam contact. If organic bioburden survives into the autoclave, steam cannot penetrate the dried coagulum; heat bakes proteins onto steel, shielding viable bacterial spores beneath.
Human Visual Inspection Gap
Under standard hospital overhead lighting, human visual inspection fails to detect up to 40% of residual micro-bioburden. Hospital-acquired surgical site infections (SSIs) cost healthcare systems over $3.3 Billion annually in avoidable morbidity.
CENTRAL ENGINEERING QUESTION
“Can a compact, low-cost optoelectronic system combining 365 nm UV-A excitation and 850 nm NIR reflectance highlight candidate bioburden regions for human verification before instruments enter the autoclave?”
Scope: Experimental pre-sterilization inspection aid. Does not replace validated washer-disinfector protocols.

## Slide 3
STERILE PROCESSING WORKFLOW
Where BloodTrace Fits in Hospital Reusable Medical Device Reprocessing
03 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
01
Point-of-Use Care
SURGICAL SUITE / BEDSIDE
Immediate wiping at point of use; application of enzymatic foam/gel to prevent coagulum from drying before transport.
02
Cleaning & Decontamination
CENTRAL STERILE DEPT (CSSD)
Ultrasonic immersion, enzymatic washing, and thermal disinfection in washer-disinfector units. Must precede sterilization.
03
BloodTrace Optical Screening
PRE-PACKAGING QUALITY GATE
3-second non-destructive screening. Flags hidden residue in hinges and serrations before instruments are wrapped in peel packs.
04
Terminal Sterilization
AUTOCLAVE CHAMBER
Pressurized steam sterilization (134°C, 30 psi). Validated only for physically clean instruments free of organic soil.
KEY COMPETITIVE ADVANTAGES OVER EXISTING HOSPITAL CLEANING TESTS
• Non-Destructive & Zero Consumables: Protein swab assays (ninhydrin, biuret) and ATP bioluminescence consume costly chemical reagents ($3–$8/test) and destroy sample traces. BloodTrace uses zero chemical consumables.
• Spatial Defect Localization: Chemical swabs only yield an aggregate pass/fail number without pinpointing location. BloodTrace maps the exact crevice, hinge, or serration requiring mechanical brush re-cleaning.
• Rapid Throughput: Operates in < 3 seconds per instrument vs. 5–10 minute incubation for swab-based enzymatic reactions.

## Slide 4
OPTICAL PHYSICS & METHODOLOGY
Dual-Wavelength Differential Spectroscopy: Overcoming Single-Band Pitfalls
04 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
365 nm UV-A Autofluorescence
EXCITATION OF NATIVE FLUOROPHORES
Excites cellular proteins, tryptophan, NADH, and porphyrins in organic bioburden, inducing secondary fluorescence emission in the 420–480 nm band. Clean passivated surgical stainless steel exhibits virtually zero autofluorescence.
850 nm NIR Absorption
HEMOGLOBIN SPECTRAL SIGNATURE
Blood exhibits intense absorption peaks in the near-infrared band (700–900 nm) due to oxyhemoglobin and methemoglobin electronic transitions. Clean electropolished steel specularly reflects ~85% of 850 nm light, creating sharp contrast.
Normalized Difference Index (NDI)
DIFFERENTIAL PIXEL SUBTRACTION
Single-band UV suffers false positives from lint, detergent residues, and rust. Cross-correlating UV emission with NIR absorption ratio isolates true biological coagulum and cancels ambient glare and surface finish variations.
NORMALIZED DIFFERENTIAL MULTI-SPECTRAL EQUATION & OPTICAL PIPELINE
Step 1: Strobe
Arduino switches 365nm UV then 850nm NIR LEDs via low-side MOSFETs.
Step 2: Capture
CMOS USB camera acquires aligned frames; BPW34 photodiode logs total flux.
Step 3: Compute NDI
NDI = (I_UV - I_NIR) / (I_UV + I_NIR + ε) computed per pixel.
Step 4: Segment
Adaptive Otsu threshold + morphological opening yields candidate contours.

## Slide 5
HARDWARE & ELECTRICAL ENGINEERING
Embedded Architecture, Circuit Interconnects & Bill of Materials
05 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
EMBEDDED HARDWARE SUBSYSTEMS
• Arduino UNO R4 Minima
Renesas RA4M1 32-bit ARM Cortex-M4 @ 48 MHz with 14-bit ADC resolution; coordinates LED microsecond strobing and telemetric USB serial streaming.
• UV Illumination Channel
High-power 365 nm UV-A LED driven via Pin D9 using AO3400 logic-level N-MOSFET (1 kΩ gate resistor, 10 kΩ pulldown to prevent floating gate during boot).
• NIR Illumination Channel
Dual 850 nm NIR LEDs driven via Pin D10 using AO3400 N-MOSFET with precision current-limiting ballast resistors (V_drop = 1.6V @ 100mA).
• Single-Point Optical Sensor
BPW34 silicon PIN photodiode coupled to high-speed transimpedance amplifier module on Analog Pin A0 (Measures aggregate reflected light flux).
• Spatial Image Acquisition
USB CMOS camera with IR-cut filter removed; sends uncompressed video frames directly to host laptop (never loaded onto microcontroller memory).
PINOUT INTERCONNECT MATRIX
Pin | Connected Hardware | Signal / Protection
D9 (PWM) | UV MOSFET Gate (AO3400) | 1 kΩ series + 10 kΩ pull
D10 (PWM) | NIR MOSFET Gate (AO3400) | 1 kΩ series + 10 kΩ pull
A0 (ADC) | BPW34 Transimpedance Out | 0–3.3V Analog Voltage
D2 (GPIO) | Tactile Trigger Pushbutton | INPUT_PULLUP to GND
5V / GND | Sensor Module / Common GND | Isolated clean ground rail
BILL OF MATERIALS (PROTOTYPE COST ESTIMATE)
• Arduino UNO R4 Minima Board: ~$20.00
• Optical LEDs (365nm UV + 850nm NIR): ~$12.00
• BPW34 Photodiode + Transimpedance Amp: ~$8.00
• USB CMOS Camera Module (IR wide-band): ~$18.00
• Enclosure, Baffles, MOSFETs, Passives: ~$6.00
• TOTAL PROTOTYPE FABRICATION COST: ~$64.00

## Slide 6
COMPUTER VISION PIPELINE
5-Stage Algorithmic Pipeline: From Optical Frames to Registered Heatmap
06 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
01. Acquisition
Synchronized multi-spectral capture. Takes visible baseline, 365 nm UV, and 850 nm NIR frames with locked exposure, white balance, and fixed camera distance.
02. Registration
Pixel-aligned spatial homography and dark-frame background subtraction. Normalizes ambient stray light and compensates for curved metallic reflections.
03. Difference
Calculates Normalized Difference Index (NDI) between UV emission and NIR absorption. Cancels specular glare and metallic machining marks.
04. Filtering
Morphological opening and Gaussian blurring remove high-frequency camera noise. Adaptive dynamic Otsu threshold isolates coagulated bioburden crusts.
05. Heatmap & ROI
Contour extraction generates spatial bounding boxes. Computes % contaminated area and assigns simulation confidence score per region.
TRIPARTITE SCREENING OUTCOME TAXONOMY
PASS / UNFLAGGED (Green)
No candidate region detected. Steel surface reflects uniform NIR (~760 a.u.) with baseline UV (<20 a.u.). Cleared for packaging.
FAIL / RESIDUE (Red)
Candidate contamination detected. Significant NIR absorption drop (>40%) and elevated UV fluorescence. Routed to ultrasonic re-wash.
AMBIGUOUS / REVIEW (Amber)
Optical variance detected (e.g. dried detergent film or water scale) without distinct hemoglobin signature. Routed to manual microscopy.

## Slide 7
EXPERIMENTAL VALIDATION & TESTING
Controlled Testing Protocol, Confounder Controls & Quantitative Metrics
07 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
FOUR CONTROLLED EXPERIMENTAL TEST GROUPS
1. Clean Reference Surface (Negative Control)
Electropolished AISI 316L stainless steel coupons; ultrasonic cleaned and autoclaved. Demonstrates zero false-positive rate under baseline settings.
2. Surrogate Blood Challenge (Positive Challenge)
Safely prepared, non-hazardous protein/hemoglobin surrogate spots (1 μL – 50 μL) dried under controlled ambient temperature. Tests spatial recovery and detection limit.
3. Detergent Residue Controls (Interference)
Surfaces dried with residual enzymatic hospital cleaning detergent. Tests whether algorithm distinguishes detergent fluorescence from true blood bioburden.
4. Water Marks & Mineral Scale (Confounders)
Distilled vs. tap water evaporated scale. Tests rejection of specular drying rings and inorganic salt deposits.
STATISTICAL EVALUATION METRICS & TARGET CRITERIA
Sensitivity (True Positive Rate)
Target: ≥ 92%
Formula: TP / (TP + FN) × 100%
Measures fraction of true contaminated bioburden regions correctly flagged by the system.
Specificity (True Negative Rate)
Target: ≥ 95%
Formula: TN / (TN + FP) × 100%
Measures fraction of clean steel surfaces correctly left unflagged without nuisance alarms.
False Positive Rate (FPR)
Target: ≤ 5%
Formula: FP / (FP + TN) × 100%
Guarantees hospital technicians are not overwhelmed by false alarms on sterile instruments.
Intra-Sample Repeatability
Target: CV ≤ 4.5%
Formula: CV = (σ / μ) × 100%
Calculates consistency across 5 consecutive scans of the identical specimen.

## Slide 8
FUTURE SCOPE & COMMERCIALIZATION
Development Roadmap: From College Proof-of-Concept to Clinical Deployment
08 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
Phase 1: Optical Hardware Refinement
NEXT 6 MONTHS (LABORATORY PROTOTYPE)
• 3D-Printed Inspection Cassette: Enclosed light-tight matte chamber with integrated mechanical instrument positioning jig.
• Narrowband Optical Filters: 850nm bandpass optical filter mounted to camera lens to eliminate ambient fluorescent lamp spillover.
• Custom PCB Fabrication: Replace breadboard with integrated STM32/ESP32-S3 PCB, onboard MOSFET gate drivers, and low-noise LDOs.
Phase 2: Algorithmic Machine Learning
MONTHS 6 – 12 (ALGORITHM & EDGE AI)
• Lightweight Edge Neural Network: Train MobileNetV4 / YOLOv8-nano on multi-hospital surgical tool imagery for joint localization.
• Multi-Alloy Spectral Library: Extend optical compensation models to matte satin, bead-blasted, and titanium nitride (TiN) gold finishes.
• Automated Box-Lock Tracking: Automatically segment high-risk anatomical crevices (box-locks, serrations, ratchets).
Phase 3: IP & Hospital Clinical Validation
MONTHS 12 – 18 (TRANSLATION & DILIGENCE)
• Institutional IP Filing: File provisional patent via MNIT Jaipur IPR cell covering dual-wavelength optical geometry & NDI algorithm.
• Hospital Trial Partnership: Pilot non-clinical bench testing with SMS Hospital / Medical College Central Sterile Processing Department (CSSD).
• Regulatory Pathway: Target FDA 510(k) Class II clearance as an accessory to medical device washer-disinfectors.

## Slide 9
REFERENCES & REGULATORY COMPLIANCE
Academic Literature, Infection Control Guidelines & Responsible Disclaimers
09 / 09
BLOODTRACE™ — MNIT JAIPUR • DEVANSH LILA • CONFIDENTIAL & PROPRIETARY
AUTHORITATIVE REGULATORY & SCIENTIFIC CITATIONS
1. U.S. Food and Drug Administration (FDA)
Reprocessing Medical Devices in Health Care Settings: Validation Methods and Labeling. Guidance for Industry and FDA Staff. — https://www.fda.gov/medical-devices/reprocessing-reusable-medical-devices/how-are-reusable-medical-devices-reprocessed
2. Centers for Disease Control and Prevention (CDC / HICPAC)
Rutala, W. A., & Weber, D. J. (2016). Disinfection and Sterilization in Healthcare Facilities: An Overview and Current Issues. Infectious Disease Clinics of North America, 30(3), 609–637. — https://www.cdc.gov/infection-control/hcp/disinfection-sterilization/healthcare-equipment.html
3. PubMed Central (PMC) Bioburden Standards
Residual protein and bioburden evaluation on reusable surgical instruments: Limits of naked-eye visual inspection. PMCID: PMC8909895 / PMC7134755. — https://pmc.ncbi.nlm.nih.gov/articles/PMC8909895/
4. Association for the Advancement of Medical Instrumentation (AAMI)
ANSI/AAMI ST79 & TIR30: Comprehensive guide to steam sterilization and sterility assurance in health care facilities. Establishing < 6.4 μg/cm² residual protein benchmarks. — https://www.aami.org/
⚠ MANDATORY CLINICAL & SAFETY DISCLAIMER FOR JUDGES AND FACULTY
• BloodTrace™ is an engineering research proof-of-concept simulation and prototype. It is NOT an FDA-cleared in-vitro diagnostic device.
• The system provides exploratory optical screening to identify candidate areas that warrant human inspection and re-cleaning. It does not replace validated chemical/thermal cleaning, disinfection, or terminal autoclaving.
• Results are reported strictly as 'Candidate Residue' or 'Suspected Contamination'. An absence of optical contrast does not prove chemical cleanliness, biological sterility, or instrument safety for patient surgical use.